California State Polytechnic University, Pomona

Environmental Health & Safety

BIOSAFETY INSPECTION CHECKLIST AND CORRECTION FORM

This form is to be used to document inspection of areas were Biosafety Level 1 and/or 2 Agents are used or stored

| Principal Investigator: | Building/Room: |
| --- | --- |
| Inspection Date: | Biosafety Level: |
| Select Agents/ Toxins: |
| Please circle all that apply: [rDNA], [Bloodborne Pathogens] Please complete additional checklists below. |

When information is entered in this form, the form is to be considered as “Sensitive Select Agent Information.”

- Biosafety Level 1 may stop after completing the BSL1 checklist.

- Biosafety Level 2 must complete BSL1 and BSL2 checklist

- Both BSL1 and BSL2 must complete the rDNA Checklist if applicable, or indicate N/A

- Both BSL1 and BSL2 must complete the Bloodborne Pathogens Checklist if applicable or indicate N/A

| Biosafety Level 1 (BSL1)Checklist: Minimum containment level required for all labs using biological agents | YES | NO | N/A |
| --- | --- | --- | --- |
| 1. Principal Investigator (PI) is familiar with the most current CDC/NIH “Biosafety in Microbiological and Biomedical Laboratories”. |  |  |  |
| 2. Access to the laboratory is limited or restricted at the discretion of the PI while experiments are in progress. The PI has the final responsibility for assessing each circumstance and determining who may enter or work in the laboratory. (BMBL: A1)(NIH:G-11-B-1-a) |  |  |  |
| 3. The PI assures proper instruction of lab staff in microbiological techniques, approved protocols, and emergency procedures, and compliance with these instructions. Documentation of training is kept by the PI (contents, trainer, attendees, and date). [7 CFR 12 (b)] |  |  |  |
| 4. A plan for emergencies has been developed and is available to lab employees. This includes fire, hazardous material spills (chemical, biological, radiological), injuries, and exposures. Spill plans must specifically address the hazards in the laboratory. [7 CFR 331.12 (a)] |  |  |  |
| 5. Laboratory coats, gowns, or uniforms are worn to prevent contamination or soiling of street clothes. This protective clothing is removed and left in the laboratory before leaving for or travel through non-laboratory areas (e.g., cafeteria, library, administrative offices, and public corridors). All protective clothing is disposed of in the laboratory, laundered by the institution, or autoclaved and laundered at home by personnel. (BMBL: C2) & (NIH: G-II-B-2f) |  |  |  |

| Biosafety Level 1 (BSL1)Checklist: (continued) | YES | NO | N/A |
| --- | --- | --- | --- |
| 6. Suitable gloves (e.g., nitrile, vinyl) should be worn at all time. (BMBL: C4) Gloves are disposed of when contaminated, removed when work is completed, and are not worn outside the laboratory. Disposable gloves are not washed or reused. Hands are washed after glove use. |  |  |  |
| 7. The laboratory is designed to easily clean and kept neat and clean. Carpet and rugs in laboratories are not permitted. (BMBL:D3)(NIH:G-11-B-3-a) |  |  |  |
| 8. Work surfaces are decontaminated daily and immediately following spills or splash of potentially infectious material with appropriate disinfectant. The decontaminant used is _________________________. (BMBL:A7) |  |  |  |
| 9. All biological wastes are decontaminated before disposal. Procedure: _____________________________________________________. (NIH:G-11-B-1-c) (NIH:G-11-B-1-b) |  |  |  |
| 10. Contaminated items are decontaminated before washing, reuse or disposal. |  |  |  |
| 11. Materials decontaminated away from the laboratory are packaged in durable leak-proof containers, which are closed before removal from the laboratory. (NIH: G-II-B-2) |  |  |  |
| 12. Lab doors must be self close and are kept closed while experiments are in progress. Closed doors facilitate directional airflow, special ventilation used to control odors or aerosols. (BMBL:D1) |  |  |  |
| 13. Eating, drinking, applying cosmetics, and food storage are not permitted in the lab. (BMBL: A3) (NIH: G-II-B1-e) |  |  |  |
| 14. Hand washing facilities are available, and persons wash their hands after handling recombinant DNA or other biohazardous materials and before leaving the laboratory.(NIH:G-II-B-1-h) |  |  |  |
| 15. A plumbed eyewash station is available within 50 feet & 10 seconds of the hazard location. The eyewash is activated weekly. (BMBL:D8) |  |  |  |
| 16. Pippetting by mouth is prohibited. Mechanical pipetting devices are used. (BMBL: A4) |  |  |  |
| 17. Care is taken to minimize the creation of aerosols. (NIH: G-II-B1-g) |  |  |  |
| 18. Cultures, stocks, contaminated plastic ware, and other non-sharps wastes are autoclaved prior to disposal. Consult specific University or College disposal requirements (e.g., clear autoclave bags, red biohazard bags). (BMBL:A8-a) |  |  |  |
| 19. Culture fluids and other contaminated liquid wastes are autoclaved or decontaminated with a suitable disinfectant before disposal down the sanitary drain. (BMBL:A8) |  |  |  |
| 20. Special containment devices or equipment such as a biological safety cabinet is generally required for manipulations of agents assigned to Biosafety Level 1. (NIH-G-11-B-3-a) |  |  |  |
| 21. The use of needles and other sharps is avoided when alternatives are available. Plasticware should substitute for glassware whenever possible. Broken glassware must not be handled directly. Instead, it must be removed using a squeegee/ scraper and dustpan, tongs, or forceps. (BMBL:A5-d) |  |  |  |

- If your work only requires Biosafety Level 1 containment, stop here. If your work requires Biosafety Level 2 containment, complete the rest of the checklist.

| Biosafety Level 2 (BSL2) Containment | YES | NO | N/A |
| --- | --- | --- | --- |
| 22. The Biosafety plan must contain sufficient information and documentation to describe the biosafety and containment procedures. [9 CFR 121.12 (a)] |  |  |  |
| 23. PI has established policies whereby only persons who are advised of the potential hazard and meet any specific entry requirements may enter lab or animal rooms. [BMBL: A1] |  |  |  |
| 24. The universal biohazard sign is posted on all lab access doors while experiments are in progress, and is posted on all units used to store organisms containing biohazardous materials. The sign includes name(s) of infectious agent(s), name and telephone numbers of responsible individuals, and any special entry requirements. (BMBL:A9)(BMBL:B1)(NIH:G-11-B-2-d) |  |  |  |
| 25. All persons working with human blood, body fluids or tissues receive annual OSHA Blood borne Pathogens Training and have been offered the HBV vaccination series. |  |  |  |
| 26. A biosafety manual or standard operating procedure is prepared or adopted. Personnel are advised of hazards and handling procedures, and are required to read and follow instructions on practices and procedures. Procedures and policies are routinely reviewed and updated. (7 CFR 331 Section 12(d)) |  |  |  |
| 27. Lab specific biosafety information is available, and personnel are required to be familiar with it. Information that should be included: general information regarding biohazardous agent, routes of disease transmission, recommended vaccinations, signs and symptoms of disease, personal protective equipment required, waste handling protocol, spill cleanup procedures for inside and outside containment equipment including centrifuges, exposure follow-up procedure, aerosol control procedure. See NIH Guidelines or CDC/NIH Biosafety in Microbiological and Biomedical Laboratories. (BMBL:B4)(NIH:G-11-B-2-m) |  |  |  |
| 28. Laboratories must have a sink for hand washing. The sink may be manually, hands-free, or automatically operated. It should be located near the exit door. (BMBL:D2) |  |  |  |
| 29. The PI has notified Environmental Health and Safety” regarding the use of Select Agents or USDA “High Consequence” these agents are stored in secured locations and doors are closed when work is in progress. |  |  |  |
| 30. The Principal Investigator has the final responsibility for the safety of all personnel within his/her laboratory. In general, persons who are at an increased risk of acquiring infection or for whom infection may be unusually hazardous (e.g., immunocompromised, immunosuppressed, pregnancy) may require special conditions or precautions before being allowed to work in the laboratory or animal rooms. If applicable, personnel should consult with their own personal physician or Student Health Services. (NIH:G-11-B-2-b)(NIH:G-11-B-2-c) |  |  |  |
| 31. Laboratory personnel receive appropriate immunizations or tests for the agents handled or potentially present in the laboratory (e.g., hepatitis B vaccine, TB skin testing). (BMBL:B2)(NIB:G-11-B-2-k) |  |  |  |

| Biosafety Level 2 (BSL2) Containment (Continued) | YES | NO | N/A |
| --- | --- | --- | --- |
| 32. The Principal Investigator provides specific training on the potential hazards associated with the work involved, the necessary precautions to prevent exposures, and the exposure evaluation procedures (e.g., symptoms of a disease) to all laboratory personnel. Personnel receive regular updates or additional training as necessary. Training is documented. (BMBL:B5)(BMBL:A11) |  |  |  |
| 33. Incident that may result in exposure to infectious materials must be immediately evaluated and treated according to procedures describe in the laboratory biosafety safety manual. All such incidents must be reported to the PI. Medical Evaluation, surveillance and treatment should be provided and appropriated records maintained. (BMBL:B8) |  |  |  |
| 34. BLS-2 laboratory worker should: not wash or reuse disposable gloves, worker should dispose of gloves with other contaminated laboratory waste. Hand washing protocols must be rigorously followed. (BMBL:C4-c) |  |  |  |
| 35. Face protection in the form of safety goggles and a mask or a chin-length face shield is worn when a splashing or spraying potential exists. Protective gear should be decontaminated or disposed of after use. (BMBL:C5) Persons who wear contact lenses in laboratories should also wear eye protection. (BMBL:C3) |  |  |  |
| 36. An insect or rodent control program is in effect. (BMBL: A10) (NIH: G-II-B-2-e) |  |  |  |
| 37. Animals and plant not associate to the experiments are excluded from the lab. (BMBL:A5-c) |  |  |  |
| 38. Needles and syringes are used only for parenteral injections and fluid aspiration from animals. Only locking or integral-type of syringes are used. Safe needle devices are required for use with blood borne pathogens and human materials. (BMBL:A5-c) |  |  |  |
| 39. Sharps including needles, razors, scalpels, contaminated broken glass and pasteur pipettes are disposed of in a sharps container. Careful manage of the material. (NIH: G-II-B-2-j)(BMBL:A5-a)(BMBL:A5-b) |  |  |  |
| 40. Extreme caution is used when handling needles and syringes to avoid autoinoculation and the generation of aerosols during use and disposal. (NIH:G-II-B-2-j) |  |  |  |
| 41. Contaminated needles are not sheared, bent, or recapped. (NIH: G-II-B-2-j) |  |  |  |
| 42. An autoclave for decontaminating infectious laboratory waste is available. (NIH: B G-II-B-4-f) |  |  |  |
| 43. Laboratory equipment should be routinely decontaminated, as well as, after spills, splashes, or other potential contamination. Equipment must be decontaminated before repair, maintenance, or removal from the laboratory. (BMBL:B7) |  |  |  |
| 44. A biohazard symbol is placed on equipment (e.g., incubators, freezers) where biohazardous materials are used or stored. Provision to assure properly safety cabinet performance and air system operation must be verified. (BMBL:D10) |  |  |  |
| 45. Special containment devices or equipment such as a biological safety cabinet is generally required for manipulations of agents assigned to Biosafety Level 2. (NIH-G-11-B-3-a) |  |  |  |

| Biosafety Level 2 (BSL2) Containment (Continued) | YES | NO | N/A |
| --- | --- | --- | --- |
| 46. Biosafety cabinets are certified annually: list cabinets below |  |  |  |
| Location: | Date last certified: |
| Certifier: | Model: |
| Manufacturer: | Serial No.: |
| Location: | Date last certified: |
| Certifier: | Model: |
| Manufacturer: | Serial No.: |

| 47. Centrifuges and microfuges are located within the laboratory. |  |  |  |
| --- | --- | --- | --- |
| 48. Chairs and furniture in the lab are made of non-fabric, easily cleanable materials. (NIH: G-II-B-4-c) |  |  |  |
| 49. A container of decontaminant is available in or near the biosafety cabinet when work is in progress. All procedures involving the manipulation of infectious material that may generate an aerosol should be conducted within a BSC or other physical containment devices. (BMBL:B10) |  |  |  |
| 50. Windows in the lab that can be opened are fitted with fly screens. (BMBL: D5) (NIH: G-II-B-4-e) |  |  |  |
| 51. A method for decontaminating all laboratory wastes should be available in the facility (e.g. autoclave, chemical disinfection, incineration, or other valid decontamination method.) (BMBL:D11) |  |  |  |
| 52. On campus transport (between laboratories, buildings) of cultures, tissues, or specimens is accomplished in closed, leak proof, break resistant containers, lined with absorbent material and labeled with the biohazard sign and contact information. Off campus transport complies with domestic (US DOT) and/or international regulations (ICAO), including required training. (BMBL:A8) |  |  |  |
| 53. Properly maintained biological safety cabinets, preferably Class II, or other appropriate physical containment devices are used whenever High concentrations or large volumes of infectious agents are handled. Such materials may be centrifuged in the open laboratory if sealed rotor heads, centrifuge safety cups, or gasket-containing centrifuge tubes are used. These rotors, safety cups, or tubes are packaged and opened only in a biological safety cabinet. (BMBL:B3)(NIH:G-II-B-3-a(2)) |  |  |  |
| 54. Studies using animals must also be reviewed and approved by the Animal Care & Use Committee prior to starting work. (BMBL:A1) |  |  |  |
| 55. A permanent record shall be maintained of the experimental use and disposal of each animal or group of animals. (NIH:Q-I:B2) |  |  |  |
| 56. A safety manual specific to the use of animals is prepared or adopted in consultation with the animal facility director and appropriate safety professionals. (BMBL:A2) |  |  |  |
| 57. Advance consideration should be given to emergency and disaster recovery plans, as a contingency for man-made or natural disasters (BMBL:A5) |  |  |  |

| Biosafety Level 2 (BSL2) Containment (Continued) | YES | NO | N/A |
| --- | --- | --- | --- |
| 58. Consideration should be given to the use of restraint devices and practices that reduce the risk of exposure during animal manipulations (e.g., physical restraint devices, chemical restraint medications, etc). Consideration should be given to the use of restraint devices and practices that reduce the risk of exposure during animal manipulations (e.g., physical restraint devices, chemical restraint medications, etc). (BMBL:B2) |  |  |  |
| 59. Materials to be decontaminated outside of the immediate areas where infectious materials and/or animals are housed or are manipulated must be placed in a durable, leak proof, covered container and secured for transport. The outer surface of the container is disinfected prior to moving materials. The transport container must contain a universal biohazard label. (BMBL:B3) |  |  |  |
| 60. Develop and implement an appropriate waste disposal program in compliance with applicable institutional, local and state requirements. Autoclaving of content prior to incineration is recommended. (BMBL: B3) |  |  |  |
| 61. Equipment, cages, and racks should be handled in manner that minimizes contamination of other areas. (BMBL:B4) |  |  |  |
| 62. Ventilation should be provided in accordance with the Guide for Care and Use of Laboratory Animals. The direction of airflow into the animal facility is inward; animal rooms should maintain inward directional airflow compared to adjoining hallways. A ducted exhaust air ventilation system is provided. Exhaust air is discharged to the outside without being recirculated to other rooms. Ventilation system design should consider the heat and high moisture load produced during the cleaning of animal rooms and the cage wash process. (BMBL:D6) |  |  |  |
| 63. Internal facility appurtenances, such as light fixtures, air ducts, and utility pipes, are arranged to minimize horizontal surface areas, to facilitate cleaning and minimize the accumulation of debris or fomites. (BMBL:D7) |  |  |  |
| 64. Floor drains must be maintained and filled with water, and/or appropriate disinfectant to prevent the migration of vermin and gases. (BMBL:D8) |  |  |  |
| 65. Cages should be autoclaved or otherwise decontaminated prior to washing. Mechanical cage washer should have a final rinse temperature of at least 180°F. The cage wash area should be designed to accommodate the use of high pressure spray systems, humidity, strong chemical disinfectants and 180°F water temperatures, during the cage/equipment cleaning process. (BMBL:D9) |  |  |  |
| 66. Illumination is adequate for all activities, avoiding reflections and glare that could impede vision. (BMBL:D10) |  |  |  |
| 67. If BSCs are present, they must be installed so that fluctuations of the room air supply and exhaust do not interfere with proper operations. HEPA filtered exhaust air from a Class II BSC can be safely re-circulated back into the laboratory environment if the cabinet is tested and certified at least annually and operated according to manufacturer’s recommendations. BSCs can also be connected to the laboratory exhaust system by either a thimble (canopy) connection or a direct (hard) connection. Correct performance of the BSCs should be recertified at least once a year. (BMBL:D11) |  |  |  |
| 68. If vacuum service (i.e., central or local) is provided, each service connection should be fitted with liquid disinfectant traps and an in-line HEPA filter, placed as near as practicable to each use point or service cock. Filters are installed to permit in-place decontamination and replacement. (BMBL:D12) |  |  |  |

| rDNA Requirements Checklist: Minimum containment level required for all labs using biological agents | YES | NO | N/A |
| --- | --- | --- | --- |
| 1. Principal Investigator (PI) is familiar with the most current “NIH Guidelines for Research Involving Recombinant DNA molecules” and CDC/NIH “Biosafety in Microbiological and Biomedical Laboratories”. |  |  |  |
| 2. The PI will report within 30 days to the IBC and NIH (ORDA) all significant problems with and violations of the NIH Guidelines and all significant research related accidents and illnesses. |  |  |  |
| 1. When an animal containing recombinant DNA or a recombinant DNA-derived organism is euthanized or dies, the carcass shall be disposed of to avoid its use as food for human beings or animals unless food use is specifically authorized by appropriate Federal Agent. (NIH:Q-I:B1) |  |  |  |
| 2. The biosafety plan must contain sufficient information and documentation to describe the biosafety and containment procedures. NIH:Q-1: B-2 |  |  |  |
| 3. The containment area shall be locked. NIH: Q-II: B-1-a-(1) |  |  |  |
| 4. The containment area shall be patrolled or monitored at frequent intervals. NIH: Q-II: B-1-a-(2) |  |  |  |
| 5. The containment building shall be controlled and have a locking access. NIH: Q-II: B-1-a-(3) |  |  |  |
| 6. The Animal Facility Director shall establish policies and procedures whereby only persons who have been advised of the potential hazard and who meet any specific entry requirements (e.g., vaccination) may enter the laboratory or animal rooms. NIH: Q-II: B-1-a-(4) |  |  |  |
| 7. Animals of the same or different species, which are not involved in the work being performed, shall not be permitted in the animal area. NIH: Q-II: B-1-a-(5) |  |  |  |
| 8. Contaminated materials that are decontaminated at a site away from the laboratory shall be placed in a closed durable leak-proof container prior to removal from the laboratory. NIH: Q-II: B-1-b-(1) |  |  |  |
| 9. Needles and syringes shall be promptly placed in a puncture-resistant container and decontaminated, preferably by autoclaving, before discard or reuse. NIH: Q-II: B-1-b-(2) |  |  |  |
| 10. When the animal research requires special provisions for entry (e.g., vaccination), a warning sign incorporating the universal biosafety symbol shall be posted on all access doors to the animal work area. The sign shall indicate: (i) the agent, (ii) the animal species, (iii) the name and telephone number of the Animal Facility Director or other responsible individual, and (iv) any special requirements for entering the laboratory. NIH: Q-II: B-1-c-(1) |  |  |  |
| 11. Laboratory coats, gowns, smocks, or uniforms shall be worn while in the animal area or attached laboratory. Before entering non-laboratory areas (e.g., cafeteria, library, administrative offices), protective clothing shall be removed and kept in the work entrance area. NIH: Q-II: B-1-d-(1) |  |  |  |

| rDNA Requirements Checklist: (Continued) | YES | NO | N/A |
| --- | --- | --- | --- |
| 12. Special care shall be taken to avoid skin contamination with microorganisms containing recombinant DNA. Impervious and/or protective gloves shall be worn when handling experimental animals and when skin contact with an infectious agent is unavoidable. NIH: Q-II: B-1-d-(2) |  |  |  |
| 13. Any incident involving spills and accidents that result in environmental release or exposures of animals or laboratory workers to organisms containing recombinant DNA molecules shall be reported immediately to the Animal Facility Director, Institutional Biosafety Committee, NIH/OBA, and other appropriate authorities (if applicable). Reports to the NIH/OBA shall be sent to the Office of Biotechnology Activities, National Institutes of Health, 6705 Rockledge Drive, Suite 750, MSC 7985, Bethesda, MD 20892-7985 (20817 for non-USPS mail), 301-496-9838, 301- 496-9839 (fax). Medical evaluation, surveillance, and treatment shall be provided as appropriate and written records maintained. If necessary, the area shall be appropriately decontaminated. NIH: Q-II: B-1-e-(1) |  |  |  |
| 14. When appropriate and giving consideration to the agent handled, baseline serum samples shall be collected and stored for animal care and other at-risk personnel. Additional serum specimens may be collected periodically depending on the agent handled and the function of the animal facility. NIH: Q-II: B-1-e-(2) |  |  |  |
| 15. Biological materials removed from the animal containment area in a viable or intact state shall be transferred to a non-breakable sealed primary container and then enclosed in a non-breakable sealed secondary container. All containers, primary and secondary, shall be disinfected before removal from the animal facility. Advance approval for transfer of material shall be obtained from the Animal Facility Director. Packages containing viable agents may only be opened in a facility having an equivalent or higher level of physical containment unless the agent is biologically inactivated or incapable of reproduction. NIH: Q-II: B-1-f-(1) |  |  |  |
| 16. All genetically engineered neonates shall be permanently marked within 72 hours after birth, if their size permits. If their size does not permit marking, their containers should be marked. In addition, transgenic animals should contain distinct and biochemically assayable DNA sequences that allow identification of transgenic animals from among non-transgenic animals. NIH: Q-II: B-1-g-(1) |  |  |  |
| 17. Needles and syringes shall be used only for parenteral injection and aspiration of fluids from laboratory animals and diaphragm bottles. Only needle-locking syringes or disposable syringe needle units (i.e., needle is integral to the syringe) shall be used for the injection or aspiration of fluids containing organisms that contain recombinant DNA. Extreme caution shall be used when handling needles and syringes to avoid autoinoculation and the generation of aerosols during use and disposal. Following use, needles shall not be bent, sheared, replaced in the needle sheath or guard, or removed from the syringe. Needles and syringes shall be promptly placed in a puncture-resistant container and decontaminated, preferably by autoclaving, before discard or reuse. NIH: Q-II: B-1-g-(2) |  |  |  |
| 18. Appropriate steps should be taken to prevent horizontal transmission or exposure of laboratory personnel. If the agent used as a vector is known to be transmitted by a particular route (e.g., arthropods), special attention should be given to preventing spread by that route. In the absence of specific knowledge of a particular route of transmission, all potential means of horizontal transmission (e.g., arthropods, contaminated bedding, or animal waste, etc.) should be prevented. NIH: Q-II: B-1-g-(3) |  |  |  |
| rDNA Requirements Checklist: (Continued) | YES | NO | N/A |
| 19. Eating, drinking, smoking, and applying cosmetics shall not be permitted in the work area. NIH: Q-II: B-1-g-(4) |  |  |  |
| 20. Individuals who handle materials and animals containing recombinant DNA molecules shall be required to wash their hands before exiting the containment area. NIH: Q-II: B-1-g-(5) |  |  |  |
| 21. A double barrier shall be provided to separate male and female animals unless reproductive studies are part of the experiment or other measures are taken to avoid reproductive transmission. Reproductive incapacitation may be used.NIH: Q-II: B-1-g-(6) |  |  |  |
| 22. The containment area shall be in accordance with state and Federal laws and animal care requirements. NIH: Q-II: B-1-g-(7) |  |  |  |
| 23. A biosafety manual shall be prepared or adopted. Personnel shall be advised of special hazards and required to read and follow instructions on practices and procedures. NIH: Q-II: B-1-g-(8) |  |  |  |
| 24. Animals shall be contained within an enclosed structure (animal room or equivalent) to minimize the possibility of theft or unintentional release and to avoid arthropod access. The special provision to avoid the entry or escape of arthropods from the animal areas may be waived if the agent in use is not known to be transmitted by arthropods. NIH: Q-II: B-2-a |  |  |  |
| 25. Surfaces shall be impervious to water and resistant to acids, alkalis, organic solvents, and moderate heat. NIH: Q-II: B-2-b |  |  |  |
| 26. The animal containment area shall be designed so that it can be easily cleaned. NIH: Q-II: B-2-c |  |  |  |
| 27. Windows that open shall be fitted with fly screens. NIH: Q-II: B-2-d |  |  |  |
| 28. An autoclave shall be available for decontamination of laboratory wastes. NIH: Q-II: B-2-e |  |  |  |
| 29. If arthropods are used in the experiment or the agent under study can be transmitted by an arthropod, interior work areas shall be appropriately screened (52 mesh). All perimeter joints and openings shall be sealed and additional arthropod control mechanisms used to minimize arthropod entry and propagation, including appropriate screening of access doors or the equivalent. NIH: Q-II: B-2-f |  |  |  |

| Bloodborne Requirements Checklist: | YES | NO | N/A |
| --- | --- | --- | --- |
| A. ECP Determination T8, CCR, 5193… |  |  |  |
| 1. Are employees exposed to body fluids? (c)(1) |  |  |  |
| 2. Is there a written ECP? (c)(1)(A) |  |  |  |
| 3. Is there a Sharps Injury Surveillance Study protocol? (c)(2)(A) 1 |  |  |  |
| 4. Is there a list of job classes where all employees have an occupational exposure? (c)(2)(A) 1 |  |  |  |
| 5. Is there a list of job classes where some employees have an occupational exposure? (c)(2)(A) 2 |  |  |  |
| 6. Is there a list of tasks and procedures where employees have an occupational exposure? (c)(2)(A) 3 |  |  |  |
| 7. Is a copy of the ECP accessible to employees? (c)(1)(C) |  |  |  |
| 8. Is the ECP reviewed and updated annually? (c)(1)(D) |  |  |  |
| B. Methods of Compliance |  |  |  |
| 1. Are Universal procedures Observed? (d)(1) |  |  |  |
| 2. Are appropriate engineering Controls used to minimize employee exposure? Including: (d)(2)(A) (d)(3)(A) |  |  |  |
| 1)___________________________________ 2) |  |  |  |
| 3)___________________________________ 4) |  |  |  |
| 5)___________________________________ 6) |  |  |  |
| 3. Are work practices and controls used to minimize employee exposure? (d)(2)(A) |  |  |  |
| 4. Are hand washing facilities readily accessible to employees? (d)(2)(C) |  |  |  |
| If not are antiseptic hand cleaners, wet wipes? (d)(2)(D) |  |  |  |
| 6. Do employees wash hands after contacting blood or OPIM? (d)(2)(F) |  |  |  |
| 7. Needles and Sharps are a) Handled appropriately according to Universal Precautions methods after being contaminated? (d)(2)(G) |  |  |  |
| b) Are reusable contaminated sharps placed in appropriate containers? (d)(2)(H) |  |  |  |
| c) Are sharps containers: Puncture resistant? (d)(2)(H) 1 |  |  |  |
| appropriately labeled and color-coded? (d)(2)(H) 2 |  |  |  |
| ii). leak proof on the sides and bottom? (d)(2)(H) 3 |  |  |  |
| Bloodborne Requirements Checklist: Minimum containment level required for all labs using biological agents | YES | NO | N/A |
| iii.) Do they provide safe access for housekeeping staff to retrieve reusable sharps? (d)(4)(B) 5 |  |  |  |
| 8. Are Food, drink, smoking, and applying cosmetics prohibited in areas where blood and OPIM are present (d)(2)(I) |  |  |  |
| 9. Are potentially infectious specimens placed in leak proof containers during handling, processing, storage, transportation or shipping? (d)(2)(M) 1 |  |  |  |
| C. Personal Protective Equipment (PPE) |  |  |  |
| 1. Is appropriate PPE available to Employees? (d)(3)(A) |  |  |  |
| 2. Is PPE used properly and consistently? (d)(3)(B) |  |  |  |
| 3. Is PPE accessible at the worksite? (d)(3)(C) |  |  |  |
| 4. Is PPE the correct size and fit properly? (d)(3)(C) |  |  |  |
| 5. Is the PPE clean and in good working order? (d)(3)(D)&(E) |  |  |  |
| 6. Is PPE stored Properly? (d)(3)(H) |  |  |  |
| 7. Gloves: |  |  |  |
| Used when exposure to blood or OPIM can be reasonably anticipated? (d)(3)(I) |  |  |  |
| Are single use gloves disposed of when barrier is compromised or after one use? (d)(3)(I) 1&2 |  |  |  |
| Are reusable utility gloves inspected and decontaminated effectively? (d)(3)(I) 3 |  |  |  |
| Are gloves used for voluntary blood drives or donations available to phlebotomists? (d)(3)(I) 4 |  |  |  |
| Required when employees have cuts on their hands? (d)(3)(I) 4 i |  |  |  |
| Required when employees are receiving training? (d)(3)(I) 4 d iii |  |  |  |
| 8. Masks, eye protection, and face shield are worn when there is reasonable anticipated splashing, spray, spatter or droplets of blood or OPIM. (d)(3)(J) |  |  |  |
| 9. Gowns, aprons, and other protective clothing when appropriate. (d)(3)(K) & (L) |  |  |  |
| D. Housekeeping |  |  |  |
| 1. Is worksite in clean and sanitary condition? (d)(4)(A) |  |  |  |
| 2. Is there a policy for prompt disinfecting of contaminated surfaces coverings and receptacles? (d)(4)(A) 1 &3 |  |  |  |
| 3. Is there policy describing correct and safe cleanup procedures for broken glassware? (d)(4)(A) 4 |  |  |  |
| 4. Are reusable contaminated sharps containers design to allow for safe removal of decontaminated sharps? (d)(4)(A) 5 |  |  |  |
| Bloodborne Requirements Checklist: (Continued) | YES | NO | N/A |
| Regulated Waste |  |  |  |
| a) Are contaminated Sharps containers closable? (d)(4)(C) 1 a i |  |  |  |
| Puncture resistant? (d)(4)(C) 1 a ii |  |  |  |
| leak proof? (d)(4)(C) 1 a iii |  |  |  |
| labeled as biohazard? (d)(4)(C) 1 a iv |  |  |  |
| easily accessible to users? (d)(4)(C) 1 b i |  |  |  |
| maintained upright throughout use? (d)(4)(C) 1 b ii |  |  |  |
| b) Are regulated waste containers closable? (d)(4)(C) 2 a i |  |  |  |
| Able to contain contents? (d)(4)(C) 2 a ii |  |  |  |
| leak proof? (d)(4)(C) 2 a ii |  |  |  |
| labeled as biohazard and color-coded? (d)(4)(C) 2 a ii |  |  |  |
| 5. b) Regulated waste (cont’d) |  |  |  |
| Closed prior to removal? (d)(4)(C) 2 a iv |  |  |  |
| Disposed of in accordance with applicable regulations? (a)(4)(C) 3 |  |  |  |
| 6. Laundry |  |  |  |
| a) Minimal handling and agitation (d)(4)(D) 1 |  |  |  |
| b) Bagged containerized at the location where it is used? (d)(4)(D) 1 |  |  |  |
| c) Labeled appropriately? (d)(4)(D) 1 b |  |  |  |
| d) Are bags or containers leak proof? (d)(4)(D) 1 c |  |  |  |
| Is contaminated laundry handled with gloves and other required PPE? (d)(4)(D) 2 |  |  |  |
| E. Hepatitis B Vaccination |  |  |  |
| 1. Is vaccination available at no cost to employees with occupational exposure? (First Aid responders who provide First Aid as a collateral duty are exempt) (f)(1)(B) 1 & (f)(1)(A) 1 |  |  |  |
| 3. Have employees been explained hepatitis “C” exposure and that vaccine is not available, (3) |  |  |  |
| F. Post-Exposure and Evaluation Follow-up |  |  |  |
| 1. When employee declines Post-exposure evaluation from the employer’s Health-Care provider, are other provisions made? (f)(1)(A) |  |  |  |
| 2. Is a Sharps Injury Log form used |  |  |  |
| 3. Are routes of entry circumstances documented? (f)(3)(A)(c)(1)(B)(4)(5) |  |  |  |

| Bloodborne Requirements Checklist: (Continued) | YES | NO | N/A |
| --- | --- | --- | --- |
| 4. Are procedures for identification and documentation of source in place? (f)(3)(B)5 |  |  |  |
| Are evaluation and follow-up procedures in place? (f)(3)(B) |  |  |  |
| 6. Are medical records kept for each employee? (f)(3)(6) |  |  |  |
| G. Labels and Signs |  |  |  |
| 1. Warning labels affixed to all containers of regulated waste including regulated waste red bags? (g)(1)(A) 1 & (g)(1)(A) 5 |  |  |  |
| 2. If labels are used: Do the warning labels include the BIOHAZARD legend (symbol)? (g)(1)(A) 2 |  |  |  |
| Are the warning labels red with lettering in contrast color? (g)(1)(A) 3 |  |  |  |
| Are the warning labels affixed to the container of regulated waste? (g)(1)(A) 4 |  |  |  |
| 3. Is contaminated equipment labeled? (g)(1)(A) 8 |  |  |  |
| 4. Are signs posted at the entrance to Bio-hazardous areas? If yes do the bear the following: (g)(1)(B) 1 |  |  |  |
| a) BIOHAZARD symbol? (g)(1)(B) 1 |  |  |  |
| b) Name of infectious agent? (g)(1)(B) 1 |  |  |  |
| c) Special requirements for entering area? (g)(1)(B) 1 |  |  |  |
| d) Name and telephone number of responsible person? (g)(1)(B) 1 |  |  |  |
| e) Is the sign bright red with black letters or Bright yellow with red letters? (g)(1)(B) 2 |  |  |  |
| H. Information and Training |  |  |  |
| 1. Is training provided to all exposed employees? (g)(2)(A) |  |  |  |
| 2. Is training provided at time of initial assignment? (g)(2)(B) 1 |  |  |  |
| 3. Is training provided at least annually thereafter? (g)(2)(B) 2 |  |  |  |
| 4. Is training provided when changes in tasks or procedures occur? (g)(2)(B) 5 |  |  |  |
| 5. Is training provided at a level appropriate for employees education level and language? (g)(2)(B) 6 |  |  |  |
| 6. Does the training include the following: |  |  |  |
| a) An accessible copy of this standard and is an explanation available? (g)(2)(G) 1 |  |  |  |
| b) A general explanation of epidemiology and symptom of bloodborne diseases? (g)(2)(G) 2 |  |  |  |
| c) Modes of transmission? (g)(2)(G) 3 |  |  |  |
| d) Explanation of and availability of the ECP? (g)(2)(G) 4 |  |  |  |

| Bloodborne Requirements Checklist: (Continued) | YES | NO | N/A |
| --- | --- | --- | --- |
| e) Potential Exposure task recognition? (g)(2)(G) 5 |  |  |  |
| f) A General Explanation of effective procedures used in the Sharps Injury Control Program (used to complete the Sharps Injury Log & surveillance study): 1) for gathering information for the Sharps Injury Log (c) (3) |  |  |  |
| 2) for periodic determination of the frequency of use of types and brands of sharps involved in exposure incidents (c) (3) |  |  |  |
| 3) for identifying currently available engineering controls, where appropriate; (c) (3) |  |  |  |
| 4) for documenting patient safety determinations made; and (c) (3) |  |  |  |
| 5) for obtaining the active involvement of employees in reviewing and updating the ECP with respect to procedures performed by employees in their respective work areas |  |  |  |
| f) Explanation and limitations of engineering controls and PPE? (g)(2)(G) 6 |  |  |  |
| g) All Phases of handling PPE? (g)(2)(G) 7 |  |  |  |
| h) Explanation of PPE selection? (g)(2)(G) 8 |  |  |  |
| i) Information on HBV vaccination? (g)(2)(G) 9 |  |  |  |
| j) Emergency actions procedures? (g)(2)(G) 10 |  |  |  |
| k) Procedures for exposure incident? (g)(2)(G) 11 |  |  |  |
| l) Procedures and employee rights for post exposure evaluation and follow-up? (g)(2)(G) 12 |  |  |  |
| m) An explanation of signs, labels and color coding? (g)(2)(G) 13 |  |  |  |
| n) An opportunity for interactive questions? (g)(2)(G) 14 |  |  |  |
| Record Keeping |  |  |  |
| 1. Are medical records kept for each exposed employee? (h)(1)(A) |  |  |  |
| 2. Do the records include the following: |  |  |  |
| a) Name and social security number of the employee? (h)(1)(B) 1 |  |  |  |
| b) A copy of the employees HBV vaccination status and related information? (h)(1)(B) 2 |  |  |  |
| c) Follow-up information when applicable? (h)(1)(B) 3 |  |  |  |
| d) Healthcare professional written opinions. e) A copy of the Sharps Injury Control Program to include the Sharps Surveillance Study and the Sharps Injury Control Log? (h)(1)(B) 4 (c) (3) |  |  |  |
| I. e) A copy of the information provided to the healthcare professional. (h)(1)(B) 5 |  |  |  |
| 3. Does the employer have procedures to ensure confidentiality? (h)(1)(C) |  |  |  |
| 4. Do training records include dates of training sessions (h)(2)(A) 1 |  |  |  |
| Bloodborne Requirements Checklist: (Continued) | YES | NO | N/A |
| A training session summary of contents? (h)(2)(A) 2 |  |  |  |
| The name and qualifications of training provider? (h)(2)(A) 3 |  |  |  |
| The names and job titles of persons trained? (h)(2)(A) 4 |  |  |  |
| 5. Are training records maintained for 3 years? (h)(3)(B) |  |  |  |
| 6. Is the Sharps Injury Log maintained for 5 years? (c) (3) |  |  |  |
| 7. Are all medical and training records available for OSHA inspection upon request? (h)(3)(C) |  |  |  |

Note: If there are any changes to the originally approved protocol, it is the responsibility of the PI to notify the Environmental Health & Safety Officer at (909)869-4697 as soon as possible to discuss these changes.

Things you should know:

- To determine if the agent you are using is Biosafety Level 1 or 2, check these links:

http://www.absa.org/index.html or http://www4.od.nih.gov/oba/

- NOTE:  Risk Group 1 is equal to Biosafety Level 1

Risk Group 2 is equal to Biosafety Level 2

- IBC Approval must be completed PRIOR to initiation of Biosafety Level 2 experiments.

- Personal protective equipment (PPE) must be provided to employees, and must be suitable for the biological agents used and the task(s) performed. PPE includes lab coats, gloves, eye/face protection (if appropriate/necessary) and any other necessary protective measures.

- A plan for containment of SHARPS should be in place if they are used (e.g., use of sharps containers; no re-capping of needles, etc).

- Changes in the proposed use of recombinant DNA materials or biohazardous agents within this project must be documented:

Use in Whole Animals (includes lab work on animal specimens):

Contact the Animal Care and Use Committee.

Use in the Laboratory ONLY (e.g., in-vitro work):

Submit Biological Agents User Application Form and Biosafety Plan (with SOP) to the EH&S Department

For assistance, contact David Patterson, EH&S; dlpatterson@cpp.edu, phone (909) 869-4698.

(Revised 08/08/2011)